Table S3.  Calibration parameters used with Hamilton Thorne IVOS computerized semen analyzer (Beverly, MA)

Parameter	Assigned Values

Apply Sort:	2

Frames acquired:	30

Frame rate:	60 Hz

Minimum contrast:	30

Minimum Cell Size:	4 pixels

Minimum Static Contrast:	15

Straightness (STR), Threshold:	50%

Low VAP Cutoff:	10.0 µm/s

Medium VAP Cutoff:	50.0 µm/s

Low VSL Cutoff:	0.0 µm/s

Head Size, Non-Motile:	13 pixels

Head Intensity, Non-Motile:	75

Static Head Size:	0.57 to 2.91

Static Head Intensity:	0.14 to 1.84

Static Elongation:	0 to 87

Slow Cells Motile:	YES

Magnification:	0.82

Video Source:	Camera

Video Frequency:	60

Bright Field;	No

Brightness for LED:	2663

Brightness for Ident:	3000

Temperature, Set:	37°C

Cell Type:	User

Cell Depth Setup:	100.0µm

Field Selection Mode:	SELECT

Indent Active:	NO

Ident Mode:	B

Integrating Time:	1 Frames

Optics

Video Gain:	Medium

Video Brightness:	2300

Video Contrast:	40

Video Sync Level:	FW Sync_Level_100

Video Vertical Sync:	75

Sort # 2

Points in Track:	16 to 100