**S2 Table: Comparison of the results obtained for 436 specimens with rapid diagnostic test (RDT) for the detection of measles-specific IgM and with measles -specific IgM indirect and capture enzyme immunoassays (EIAs) or with reverse transcription-polymerase chain reaction (RT-qPCR) for viral detection (indeterminates as positives)**

|  |  |  |
| --- | --- | --- |
|  | **RDT with capillary blood** | **Total (missing)** |
| **Measles IgM EIA Virion/Serion or RT-qPCR** | **Positive** | **Negative** |  |
| **Positive** | 27 | 17 | 44 (1) |
| **Negative** | 18 | 368 | 386 (5) |
| **Total** | 45 | 385 | 430 (6) |

|  |  |  |
| --- | --- | --- |
|  | **RDT with oral fluid** | **Total (missing)** |
| **Measles IgM EIA Virion/Serion or RT-qPCR** | **Positive** | **Negative** |  |
| **Positive** | 20 | 23 | 43 (2) |
| **Negative** | 30 | 345 | 375 (16) |
| **Total** | 50 | 368 | 418 (18) |

|  |  |  |
| --- | --- | --- |
|  | **RDT with capillary blood** | **Total (missing)** |
| **Measles IgM EIA Virion/Serion and Measles IgM EIA Microimmune or RT-qPCR** | **Positive** | **Negative** |  |
| **Positive** | 22 | 8 | 30 (0) |
| **Negative** | 23 | 377 | 400 (6) |
| **Total** | 45 | 385 | 430 (6) |

|  |  |  |
| --- | --- | --- |
|  | **RDT with oral fluid** | **Total (missing)** |
| **Measles IgM EIA Virion/Serion and Measles IgM EIA Microimmune or RT-qPCR** | **Positive** | **Negative** |  |
| **Positive** | 16 | 13 | 29 (1) |
| **Negative** | 34 | 355 | 389 (17) |
| **Total** | 50 | 368 | 418 (18) |